Tender Notice & Documents No 312/2016

for

Supply and Installation of Scientific Equipments for Different Departmental Laboratories of the Government College University Faisalabad (GCUF)

Last Date for Submission:				20-04-2016 at 11:00 AM

Bid Opening Date:					20-04-2016 at 11:30 AM

GOVERNMENT COLLEGE UNIVERSITY FAISALABAD

#### Postal Address:

#### Government College University Faisalabad (GCUF), Allama Iqbal Road, Faisalabad

(Phone: 041 9201030 & 9201468)

GOVERNMENT COLLEGE UNIVERSITY FAISALABAD

TENDER NOTICE NO. 312 /2016

Government College University, Faisalabad (GCUF) invites sealed tenders / bids from the well reputed firms / importers / suppliers, registered with Sales Tax and Income Tax Departments, for procurement, supply and installation of “Scientific Equipments for Different Departmental Laboratories of Government College University Faisalabad (GCUF)”. Detail and Specifications of required equipments are attached in the Tender Documents.

- Tender Documents are available for download at GCUF's website www.gcuf.edu.pk and PPRA's website  www.ppra.punjab.gov.pk. The procurement shall be completed in accordance with Punjab Procurement Rules 2014, on Single Stage - Two Envelope Bidding Procedure.

- A single package containing two envelopes, Technical and Financial separate Bids, duly completed, signed, stamped, sealed and in complete conformity with Tender Documents should be dropped in Tender Box placed at Procurement Department of the GCUF, not later than 1100 Hrs on 20.04.2016 . Technical Bids will be opened first on 20.04.2016 at 1130 Hrs, and Financial Bids will be opened (as per schedule issued later on) for only those Successful Technical Bids which are acceptable to the "Technical Committee/s". Whereas, Financial Bids of unsuccessful Bidders will be returned un-opened.

- All bids must be accompanied by Tender Fee Rs; 1,000/- and Bid Security (Refundable) at the rate of 2% of the estimated price of Rs140.000 (M) in shape of CDRs in favor of "Treasurer GCUF". The CDRs of Tender Fee and Bid Security (in original) must be attached with Technical Bid & Financial Bid, respectively. Bids which are incomplete, not signed,  not on letter pad,  not sealed, without tender fee & bid security, late or submitted by other than specified mode will not be considered.

- The University may reject all bids/proposals at any time prior to the acceptance of a bid or proposal as per PPRA Punjab rule 35, however, upon bidder request, the grounds for rejection will be communicated to the concerned but no justification will be given as per PPRA rule 35 (2).

- In case of closed/forced holidays, tender opening time/date will be considered as the next working day.

Officer Incharge
Directorate of Procurement and Inventory Control
Allama Iqbal Road Faisalabad, Phone: 041-9201468 & 9201030

GOVERNMENT COLLEGE UNIVERSITY FAISALABAD

TENDER DOCUMENTS NO. 312 /2016

SUPPLY AND INSTALLATION OF SCIENTIFIC EQUIPMENTS FOR DIFFERENT DEPARTMENTAL LABORATORIES OF THE GOVERNMENT COLLEGE UNIVERSITY FAISALABAD (GCUF).

Purchase Price: Rs. 1,000/-

Government College University, Faisalabad (GCUF) invites sealed tenders / bids from the well reputed firms / importers / suppliers, registered with Sales Tax and Income Tax Departments, for procurement, supply and installation of “Scientific Equipments for Different Departmental Laboratories of Government College University Faisalabad (GCUF)”. Details & Specifications are separately attached as Appendix “A”.

GENERAL TERMS & CONDITIONS:

- Tender Opening Date & Procedure:

  - The procurement shall be completed in accordance with Punjab

Procurement Rules 2014, on Single Stage - Two Envelope Bidding

Procedure.

  - A single package containing two envelopes, Technical and Financial

Separate Bids in  complete conformity with Tender Documents will be

dropped in Tender Box placed at Procurement Department of the GCUF,

not later than 1100 Hrs on 20.04. 2016 .

  - Technical Bids will be opened first on 20.04.2016 at 1130 Hrs for

evaluation by the " GCUF Technical Committee".

  - Financial Bids of Successful Technical Bidders will be opened as per

schedule issued later on.

1.5  Financial Bids of unsuccessful Bidders will be returned un-opened.

- Tender Fee, Bid Security and Performance Security:

2.1 Technical Bid must be accompanied by Tender Fee of Rs:1,000/- in shape of

Call Deposit Receipt (CDR) in original.

2.2 Financial Bid must be accompanied by Bid Security @ 2% of the estimated /

quoted price (refundable) in shape of Call Deposit Receipt (CDR) in original.

  - CDRs must be in favor of Treasurer, Government College University, Faisalabad.

  - Performance security  @ 5% of the total bill shall be retained by the University and  refundable after expiry of the warranty period.

- Validity of Offers:

3.1 Offers shall be valid for at least 180 days from the date of submission of bids.

3.2 Withdrawal / modification of the original offer within the validity period shall

entitle the University to forfeit Bid Security.

3.3 The Competent Authority reserves the right to delete or decrease quantities.

- Failures and Terminations:

No offer of a supplier / firm will be considered if:-

  - Received without 2% Bid Security’s  Pay order / Bank Draft / CDR or less.

  - Bid received not in accordance with specifications of Tender Documents.

  - Bid received later than the date and time fixed for tender opening.

  - Bid / offer incomplete in any respect or unsigned.

  - Bid / offer ambiguous and / or conditional.

  - Bid / offer from a firm which is black listed at any level.

  - Any erasing / cutting / overwriting etc is observed.

  - GCUF reserve the right to reject  all bids at any time prior to the acceptance of  bids, as provided under Rule-35 of Punjab Procurement Rules, 2014.

  - The vendor fails to deliver the consignment within specified delivery period    strictly in accordance with the terms and conditions as laid down in the contract / supply order.

  - If situation warranted, then GCUF is authorized to forfeit the bid Security  and the firm may also be black listed.

  - The decision of the Vice Chancellor of Government College University, Faisalabad, would be final & binding on both the parties and not challengeable in any court of law.

- Scope of Work:

Following requirements define the scope of work of this Tender:

  - The vendors will be responsible for the supply, installation, commissioning of scientific equipments alongwith on campus training.

  - The vendors must ensure that the supplied equipment is fully operational, new, perform properly and meet the Technical Specification.

  - The venders shall provide comprehensive documentation of system deployed including diagrams,  configuration and manuals etc at the time of installation and commissioning.

  - The Vendors shall be responsible of providing services  for operation and maintenance of the equipment provided, without any extra cost or hidden charges during warranty period.

  - The venders shall provide any up-gradation of the software without any extra cost or hidden charges during warranty period.

- Mandatory Requirements:

  - For acceptance of bids for equipments, Bidders shall meet the Mandatory Requirements and Technical Specifications.

  - The Bidders should read carefully “Lab Equipment Description and Specification” before preparing Technical and Financial Proposals.

  - The Bidders shall provide Company profile, Authorization with principal firm(s), company experience in related field, project completed, major clients’ list, equipment’s technical details (brochures, etc.) and “Technical Requirements” by clearly mentioning make, model, origin, part number, manufacturer’s authorization and manufacturing year etc.

  - The Bidders are required to submit their technical and financial proposals for supply of each equipment separately.

  - The Bidder shall quote single option of each equipment which may be equivalent or higher to the required specifications along with bid security equal to 2% of the quoted price of equipment. No alternative/multiple options will be accepted.

  - The Bidders are required to submit their financial proposals in  foreign currency i-e US Dollar ($) only on the basis of Cost,  Insurance & Freight (CIF) value inclusive of all clearance charges, L/C charges, local carriage charges etc till delivery of equipment/s at GCUF Campuses. Bank draft/ Payorder/ CDR equal to 2% of the Total Bid Value or Total Estimated / Quoted Price should accompany the bid  as part of financial proposal as bid security drawn in favor of “Treasurer GCUF”. The bid shall not be considered without bid security.

  - The Successful Bidder shall open “Letter of Credit” at his own in favor of GCUF and letters for custom duty / excise duty / salex tax / other surcharge exemption shall be issued, wherever applicable.

  - The bidder may quote price of one item or more than one item at his own choice.

  - Quoted equipments must be of European Economic countries/ USA/ Canada/ Japan origin or equivalent.

- Terms of Payment:

  - No payment shall be made in advance to the successful bidder.

  - Payment will be made in Pak Rupees (equivalent to the exchange rate of US $ applicable at the time of supply order) through crossed cheque within 30 days after timely delivery, inspection and installation of equipments as per GCUF Rules.

  - The vender is responsible to provide import documents (B/L, shipping bill, custom documents, packing list etc.) along with invoice and delivery challan for timely release of payment.

  - The performance security @ 5%  shall be deducted from the final bill of the successful bidder and will be released after satisfactory performance of the equipments and termination of warranty period.

- Liquidated Damages:

In case of delay, GCUF reserves the right to impose a penalty not exceeding 10% of the total amount of the supply order at the rate of 2% per month of the bill or amount of delayed equipment.

Officer Incharge
Directorate of Procurement and Inventory Control
Allama Iqbal Road Faisalabad, Phone: 041-9201468 & 9201030

TENDER DOCUMENTS NO. 312 /2016

SUPPLY AND INSTALLATION OF SCIENTIFIC EQUIPMENT FOR DIFFERENTAL DEPARTMENTAL LABORATORIES OF THE GOVERNMENT COLLEGE UNIVERSITY FAISALABAD (GCUF)

| A: Bill of Cost Equipments |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sr. # | Description | Unit | Qty | Unit Price including Taxes | Total Price |
|  | Digital Oven Precise Drying | No. | 3 |  |  |
|  | Rotary Evaporator System | No. | 3 |  |  |
|  | UV/Vis Dual Beam Spectrophotometer | No. | 3 |  |  |
|  | FTIR Spectrometer (Bench-Top) | No. | 1 |  |  |
|  | Walk-in Cold Room 10'x10'xS10' | No. | 1 |  |  |
|  | Camera Fitted Light Microscope with Monitors and Accessories | No. | 1 |  |  |
|  | Inverted Microscope with Monitors and Accessories | No. | 1 |  |  |
|  | Shaker Incubator | No. | 2 |  |  |
|  | Glass Distillation Unit, Complete with Mountable Accessories | No. | 3 |  |  |
|  | Electronic/Analytical Balance (Four Digit) | No. | 10 |  |  |
|  | Shaking Water Bath (Temperature Controlled) | No. | 2 |  |  |
|  | Homogenizer | No. | 2 |  |  |
|  | Vortex Mixer | No. | 3 |  |  |
|  | Ice Flake Maker | No. | 1 |  |  |
|  | DO Meter | No. | 2 |  |  |
|  | pH Meter (Digital Bench Model) | No. | 3 |  |  |
|  | EC Meter / Conductivity Meter | No. | 3 |  |  |
|  | Vacuum Pump | No. | 3 |  |  |
|  | UV Lamp | No. | 1 |  |  |
|  | Drying Cabinet | No. | 1 |  |  |
|  | Precision/Micro Pipettes (0.5, 10, 20, 100, 200, 1000µl) | No. | 2 |  |  |
|  | Freezer -20C | No. | 1 |  |  |
|  | Blenders | No. | 2 |  |  |
|  | Liquid Nitrogen Flasks | No. | 1 |  |  |
|  | Microwave Oven | No. | 2 |  |  |
|  | Microtome with Accessories | No. | 1 |  |  |
|  | Incubator | No. | 3 |  |  |
|  | 4 Ft Laboratory Fume Hood Bench Top Model | No. | 3 |  |  |
|  | Hot Plate / Magnetic Stirrer | No. | 4 |  |  |
|  | Pipette Aid Automatic | No. | 1 |  |  |
|  | Protein Electrophoresis Apparatus with All Accessories and Power Supply | No. | 2 |  |  |
|  | Bench Top Orbital Shaker | No. | 1 |  |  |
|  | Gel Drying System | No. | 1 |  |  |
|  | Eye Wash Station | No. | 2 |  |  |
|  | Gel Documentation - Imaging System | No. | 2 |  |  |
|  | Trays for Gels | No. | 3 |  |  |
|  | Micro Kjeldhal Apparatus | No. | 1 |  |  |
|  | Electrically Heated Muffle Furnace | No. | 1 |  |  |
|  | Liquid Nitrogen Cylinder | No. | 1 |  |  |
|  | Sieve Shaker Mechanical with Sieves Set ASTM | No. | 1 |  |  |
|  | Incubator for BOD | No. | 1 |  |  |
|  | Biofilter for Aquarium | No. | 1 |  |  |
|  | Gas Flow Meter | No. | 3 |  |  |
|  | Laser Power Meter | No. | 2 |  |  |
|  | LCR Meter | No. | 2 |  |  |
|  | Constant Current and Voltage Programable Power Supply | No. | 2 |  |  |
|  | Cryostat for Liquid Nitrogen with Sample Holder | No. | 2 |  |  |
|  | Temperature Controller | No. | 2 |  |  |
|  | Oscilloscope (dual trace 4 channel) with GPIB Interface and Software | No. | 2 |  |  |
|  | Nd YAG laser (solid state laser with first second and third harmonic generations) YAG Laser with Amplifier | No. | 1 |  |  |
|  | Spectrometer | No. | 1 |  |  |
|  | Delays | No. | 1 |  |  |
|  | Optics for Laser | No. | 1 |  |  |
|  | High Temperature Furnace with Compatible Crucible Box Type (1600 °C) | No. | 1 |  |  |
|  | Ball Mill, Dry and Wet Grinding | No. | 1 |  |  |
|  | Air Sampler (Basic) | No. | 1 |  |  |
|  | Hydraulic Press | No. | 1 |  |  |
|  | Differential Scanning Calorimeter (DSC) | No. | 1 |  |  |
|  | GC-MS | No. | 1 |  |  |
|  | Digital Colony Counter | No. | 1 |  |  |
|  | Staining Rack | No. | 1 |  |  |
|  | Phase Contrast Microscope with Accessories | No. | 1 |  |  |
|  | Refrigerator | No. | 2 |  |  |
|  | Autoclave | No. | 2 |  |  |
|  | Centrifuge | No. | 1 |  |  |
|  | Horizontal Electrophoresis Apparatus | No. | 1 |  |  |
|  | Inverted Microscope With Camera | No. | 2 |  |  |
|  | Shaker Digital | No. | 1 |  |  |
|  | Microbiological Safety Cabinet Class II | No. | 2 |  |  |
|  | DIAREADER / Microplate Reader with Computer, Software and Accessories | No. | 2 |  |  |
|  | Microplate Washer | No. | 1 |  |  |
|  | Fluorescence Microscope with Trinocular Head | No. | 1 |  |  |
|  | Laminar Air Flow | No. | 2 |  |  |
|  | Lyophilizer / Freezer Dryer with all Accessories | No. | 2 |  |  |
|  | Refrigerated Centrifuge Machine, Temperature Control | No. | 1 |  |  |
|  | Thermal Cycler | No. | 1 |  |  |
|  | Sonicator | No. | 2 |  |  |
|  | Ultra Low Temperature Freezer | No. | 1 |  |  |
|  | Ordinary Light Microscope with All Acessories: | No. | 4 |  |  |
|  | Micro Dispenser | No. | 1 |  |  |
|  | Elisa Plate Reader | No. | 1 |  |  |
|  | Power Lab with Accessories | No. | 1 |  |  |
|  | Blood Autoanalyzer with Washing Solutions and Accessories | No. | 1 |  |  |
|  | Multimeter | No. | 5 |  |  |
|  | Sulphur Dioxide (SO2) Meter | No. | 4 |  |  |
|  | Oxides of Nitrogen as NO Meter | No. | 4 |  |  |
|  | Oxides of Nitrogen as NO2 Meter | No. | 4 |  |  |
|  | Tropospheric Ozone (O3) Meter | No. | 4 |  |  |
|  | Suspended Particulate Matter (SPM) Meter | No. | 4 |  |  |
|  | Respirable Particulate Matter PM10 Meter | No. | 4 |  |  |
|  | Digital CO Meter | No. | 4 |  |  |
|  | Portable Airborne Lead & Lead Dust Monitor | No. | 4 |  |  |
|  | Portable CO2 Meter | No. | 4 |  |  |
|  | Digital Sound Level Meter | No. | 4 |  |  |
|  | Ultra Micro Balance | No. | 4 |  |  |
|  | Penetrometer Micrometric Vertical Adjustment | No. | 3 |  |  |
|  | Pocket Penetrometer | No. | 5 |  |  |
|  | Stack/Flue Gas Analyzer | No. | 2 |  |  |
|  | Total Organic Carbon (TOC) Analyzer | No. | 1 |  |  |
|  | Atomic Absorption Spectrometer | No. | 1 |  |  |
|  | Ion Chromatograph | No. | 1 |  |  |
|  | Portable Total Suspended Solids (TSS) Meter | No. | 3 |  |  |
|  | Chemical Oxygen Demand (COD) Meter | No. | 1 |  |  |
|  | Biological Oxygen Demand (BOD) Meter | No. | 1 |  |  |
|  | Fully Automatic Hematology Analyzer | No. | 1 |  |  |
|  | Calorimeter: Automatic Isoperibol Calorimeter with Oxygen Combustion Bomb, Oxygen Regulator, Nitrogen Regulator, Rinse Tank, 500 Firings Kit, Printer and Ribbon | No. | 1 |  |  |
|  | Viscometer | No. | 2 |  |  |
|  | Amylogram/ Amylograph | No. | 1 |  |  |
|  | Mixogram | No. | 1 |  |  |
|  | Farinogram/ Farinograph | No. | 1 |  |  |
|  | FPLC (Fast Protein Liquid Chromatography) | No. | 1 |  |  |
|  | Texture Analyzer | No. | 1 |  |  |
|  | Fluorescence Detector | No. | 1 |  |  |
|  | Ultrasound Assisted Extraction with Accessories | No. | 1 |  |  |
|  | UHT Milk Production Unit | No. | 1 |  |  |
|  | Pilot Plant for Juice Concentrates | No. | 1 |  |  |
|  | Blanching and Washing Unit | No. | 1 |  |  |
|  | Semi-Automatic Food Packaging Machine | No. | 1 |  |  |
|  | Digital Flame Photometer | No. | 2 |  |  |

Lab Equipment Description and Specification

| - Equipment |  |  |
| --- | --- | --- |
| Sr. # | Specifications | Qty |
|  | Digital Oven Precise Drying: Adaptive Multifunctional Digital PID-microprocessor Controller with High-Definition TFT-Color Displays • Temperature range 5°C above ambient up to +300 °C • Single DISPLAY • natural convection • adjustment of pre-heated fresh air admixture by air flap control • self-diagnostics for fault analysis • Drying, High accurate Temperature. • Digital timer adjustable from 1 minute to 99 days, 23 hours • Multiple Overtemperature Protection with audible and visual alarm • Interior Stainless Steel & Textured Stainless Steel Casing Operated on 220volts, Capacity 74 Lit, interior dimension: 400 x 560 x 330 mm with two stainless steel trays. Inner size = 350 x 350 x 400 mm. | 3 |
|  | Rotary Evaporator System: Technical Specifications: - Hand / Motor lift Rotation speed (rpm) 20 – 280 Rotation speed setting Scale Drive Brushless DC motor with electronic speed control Heating capacity (W) 1,300 Temp. range heating bath (°C) 20 – 210 Temperature range accuracy (°C) ±1 Overheat protection bath (°C) Cut-off at 5°C over set temperature Bath temperature setting Digital Heating bath temperature control Electronic / Digital Secondary overtemp. Cut off (°C) 250 Material heating bath (mm) Stainless steel 316L Volume heating bath (I) 4.5Liters Supply with 1Litre evaporating and Receiving flask as standard. With Glassware set operated on 220 Volts. Accessories a. Water Aspirator pump operates efficiently on both high and low water pressure. Nickel-plated 3/8 in. NPTM water inlet. Ball type, positive-acting valve in suction tube. Vacuum better than 100 mm Hg (abs) b. Chiller Specifications: Small recirculating chiller that fulfills the chilling requirements pertinent to rotary evaporator. Cooling temperature 0 - 10 °C. Perfect compatibility for use with Rotary Evaporator. | 3 |
|  | UV/Vis Dual Beam Spectrophotometer with fixed spectral resolution: Mode: Energy, Absorption, Transmission, Reflectance, Concentration, Factor Sample Compartment dimensions (W x H x D): 364 x 185 x 260 mm Wavelength range: 190—1100 nm Spectral bandwidth: Fix 1.4 nm Scanning speed: 12000 nm/min Wavelength accuracy (Deuterium line at 656nm): ± 0.1 nm Wavelength reproducibility (with Holmium oxide filter): ≤0.02 nm Photometric reproducibility: ≤ 0.0005 A Stray light: ≤ 0.03 % T Baseline stability at 500 nm: ≤ 0.0001 (RMS) Longterm stability a 500 nm: ±0.0005 A/h Scanning speed: 12000 nm/min Auto zero function, Complete with application software, Computer, 4-8 Auto Cell Changer, Quartz cuvettes (3 ml, 1-1.5 ml) one pair each, Extra lamps etc. C = concentration & F = factor, auto-zero. Power supply may be switched: 100-220 V. Operated by panel controls or through a PC using the application software. | 3 |
|  | FTIR Spectrometer (Bench-Top): All spectrometer components, automated control of specification by periodically performed test measurements, automated selection of appropriate measurement parameters according to individual sampling accessories, status lamp in hardware and software, integrated and automated instrument tests and validation according to cGMP (OQ, PQ). Standard libraries (e.g. Pharmaceutical library), Transmission accessory & Liquid cells along with complete accessories. Outstanding signal stability and performance. Spectral range: 6,000 to 500 cm-1 with Zn Se beamsplitter Extension for operation in regions with high humidity Resolution: Better than 0.4 cm-1 Wave number accuracy: Better than 0.01 cm-1 @ 2,000 cm-1 Photometric accuracy: Better than 0.1 % T Signal-to-Noise: 5 Sec: >6,000:1 (= <7.2 x 10-5 AU noise) peak-to-peak, 4 cm-1 resolution Signal-to-Noise (achievable): 5 Sec: >8,000:1 (= <5.4 x 10-5 AU noise) peak-to-peak, 4 cm-1 resolution 1 Min: >45,000:1 (= <9.7 x 10-6 AU noise) peak-to-peak, 4 cm-1 resolution Interferometer: RockSolid , permanent aligned, High stability Optics: Gold coated mirrors Laser: Diode laser with long life time IR-source: Global,electronically stabilized for long life Detector: DigiTect detector system, high sensitivity room temperature, Scan speed: 6 velocities, 1.4 - 51 mm/sec opd Spectra rate: Up to max, 25 spectra per second @16 cm-1 optical resolution A/D converter: True 24-bit dynamic range for all scan velocities; Dual channel data acquisition Validation: VU internal validation unit, fully automated OQ/PQ test and protocols Aperture wheel: 11 positions, fixed diameters, Ranging from 250 µm to 6 mm Spectroscopy software: OPUS or equivalent: easy-to-use, full GLP compliant, full 21CFR part 11 compliant System Status Indicators Indicator lights report scan, laser, power, and source status continuously on software Instrument Qualification. IQ, OQ, PQ & DQ Integrated Platinum-ATR-accessory, Diamond, 1 bounce, QuickLock Specifications: - Diamond crystal, single reflection, edged in tungsten carbide - Work disk: stainless steel - Spectral range: 10-8,000cm-1 - Ergonomic one-finger clamp mechanism - Pressure control - Working distance (max. sample height) >20mm - Easy cleaning due to the rotatable pressure applicator - Exchangeable crystal plates (Diamond, Ge) - Purgeable and sealed beam path - Sample compartment cover adapted for Platinum-ATR-accessory. | 1 |
|  | Walk-in Cold Room 10'x10'xS10': Thicknesses of panels (150mm) 20cm modulation, for 60-85-105mm and 40cm for 150mm Floor panel finishes (phenolic resin), rust-proof Possibility of interior division or of joined cold rooms Hinged/sliding doors Maximum integration of the components to facilitate cleaning all accessories that help to comply with the current hygiene regulations should be included Storage - refrigeration unit. Temperature range -2/+5 C, Air condensing units, evaporators and controllers. | 1 |
|  | Camera Fitted Light Microscope with Monitors and Accessories: With plan achromatic objectives with long working distance (no cover glass) and wide-field eyepieces, can get clear pictures and wide view field. Coaxial coarse/fine focus system, with tensional adjustable and up stop, minimum division of fine focusing: 2μm. With vertical illuminator and transmitted illuminator, can observe surfaces of non-transparent object or transparent object. 6V 20W halogen lamp, brightness control. Trinocular, can switch to observe normally or to observe the Polarize, can send 100% of light to the binocular eyepieces or to the top port. Specifications 1. Eyepiece: Wide field WF10X(Ф18mm) and WF16X(Ф11mm) Plan dividing 10X(0.10mm/ div) 2. Objective: Plan achromatic (No cover glass): PL4X/0.10 PL5/0.12 PLL10X/0.25 PLL40X/0.65(spring) PLL60X/0.85(spring) PL100X/1.25oil(spring) PL100X/0.85dry(spring) 3. Eyepiece tube: Trinocular, Compensation(inclination of 30°), including analyzer 4. Nosepiece: Backward inner locating on ball bearing, Quintuple 5. Focus system : Coaxial coarse/fine focus system, with tensional adjustable and up stop, it is up and down adjustable area with 35mm at coaxial coarse/fine focus system, it can be used to test specimens the maximum only 20mm height, minimum division of fine focusing: 2μm. 6. Illumination unit: Reflected & Transmitted light illuminator 6V 20W, adjustable halogen lamp, 6V 20w halogen lamp, brightness control, Vertical illumination with field diaphragm, aperture diaphragm and polarizer,(Y,B,G)filter and frosted filter Abbe condenser NA.1.25 Rack & pinion adjustable, Blue filter and frosted filter Collector for halogen lamp 6V 20w halogen lamp, brightness control with polarizer Rack & pinion adjustable 7. Stage: Double layer mechanical (Size: 185mmX140mm, moving range: 75mmX50mm) 8. CCD camera unit: Standard Outfit 1. Microscope main body: 1unit 2. Wide field eyepiece 10X: 2pc 3. Plan achromatic objective 5X，10X，40X，60X: 1unit each 4. Trinocular tube (including analyzer): 1unit 5. Quadruple nosepiece: 1pc 6. 6V 20W halogen lamp: 1pc Accessories 1. Wide field eyepiece 16X (φ11mm) and Dividing 10X(φ18mm) 0.1mm/Div 2. Plan achromatic objective 4X、20X、50X、80X、100X/0.85(dry) and 100X/1.25(oil) 3. Quintuple nosepiece: 1pc 4. Filter Green, Blue and Yellow 5. Photo unit, CCD camera unit and accessories, Color TV 6. CCD Adapter 0.4 and 0.5X 7. Digital Camera CANON / NIKON. | 1 |
|  | Inverted Microscope with Monitors and Accessories: Ti-E: eyepiece 100%, left 100%, right 100%, eyepiece 20%/left 80% Ti-E/B: eyepiece 100%, left 100%, right 100%, bottom 100% Motorized port switching Two ports (tube base unit with side port, back port) can be added optionally. Via motorized nosepiece up/down movement, Stroke (motorized): up 7.5mm, down 2.5mm, Motorized (pulse motor Minimum step: 0.025µm, Maximum speed: 2.5mm/sec or higher, Motorized escape and refocus mechanism (coarse) Coarse/fine switchable, Intermediate magnification 1.5x, Light intensity control, Light on/off switch,VPD on front of body, Operation with controller, TI-TD Binocular Tube D, TI-TS Binocular Tube S, TI-TERG Ergonomic Tube, TI-T-B Eyepiece Tube Base Unit, TI-T-BPH Eyepiece Tube Base Unit for PH, TI-T-BS Eyepiece Tube Base Unit with Side Por, CFI 10x, 12.5x, 15x, TI-DS Diascopic Illumination Pillar 30W, TI-DH Diascopic Illumination Pillar 100W TI-S-ER Motorized Stage with Encoders, TI-S-E Motorized Stage—Cross travel: X110 x Y75mm, Size: W400 x D300mm (except extrusions). TI-SR Rectangular Mechanical Stage—Cross travel: X70 x Y50mm, Size: W310 x D300mm, TI-SP Plain Stage—Size: W260 x D300mm TI-SAM Attachable Mechanical Stage—Cross travel: X126 x Y84mm when used with TI-SP Plain Stage Sextuple fluorescence filter cube rotating turret, Filter cubes with noise terminator mechanism, Field diaphragm centerable, 33mm ND4/ND8 filters, 25mm heat absorbing filter Option: Motorized sextuple fluorescence filter cube rotating turret, Motorized excitation filter wheel, Motorized barrier filter wheelContrast control: Senarmont method (by rotating polarizer) Objective side prism: for individual objectives (installed in nosepiece) Condenser side prism: LWD N1/N2/NR (Dry), HNA N2/NR (Dry/Oil) types. Hot stage should be included. | 1 |
|  | Shaker Incubator: Built in timer to automatically shut off, heating system, Program incubator with built in RS 232 interface and included software, inner glass door, Electronic PID control, 3 Digit Colored LEDs display. Capacity 3.8 Cu ft, shelf: 5 Power: 220 V, 900 Watts, Temperature: 30 to 70 C, Accuracy: +/- 0.1 C, Uniformity: +/- 0.5 C, Resolution: 0.1 C, 10 to 300 rpm speed. | 2 |
|  | Glass Distillation Unit, Complete with Mountable Accessories: Output: 20 ltr/hr Electrical Power 220/240 Volts, 50 - 60 Hz, Single Phase, 3 KW. Silica heater, Cooling water required 1-2 lit/min Min. Pressure 3 psi, Weight 8-16 kgs. Conductivity 4 μs/cm, PH 5.5 to 6.5. | 3 |
|  | Electronic/Analytical Balance (Four Digit): Digital, Capacity upto 250 grams Readability: 0.001 mg, Linearity: ±0.002 mg Calibration: External Digital, Display (cm) LCD pan Size (cm): 10 diameter. 220 V, 50/60 Hz. Extra Feature: Overload Protection Features: - Automatic calibration by pressing keys with internal calibration weight - Easy setting, expanded piece counting function comparator function, formulation modes - Windows direct communication function, send balance data to excel without any data software with optional RS232C interface. | 10 |
|  | Shaking Water Bath (Temperature Controlled): Dimensions: Bath: 500 x 295 x 150mm WxDxH , Ext. 536 x 335 x 260 mm WxDxH Heater 2 kW, Capacity: 22 liters Temp. Range: Ambient +5°C – 100°C, Accuracy: ±0.1°C Microprocessor control with digital LED display and shaker control Adjustable shaking speed from 30 to 200 (OPM) oscillations/min. Over-temperature protection function Rugged stainless-steel chamber with easy to clean epoxy-powder coated exterior  Push-button convenience for entry of temperature and shaking speed Automatic motor shut-down capability, Shaking Tray, Rubber Duck, 220 V, 50/60 Hz. | 2 |
|  | Homogenizer: Particle reduction down to 1 µm. Operating range up to 2.5 liters. With Integral variable electronic speed control. Maximum speed 24000 rev/min. Overall 75 x 60 x 250 mm. For 220-240 V, 50/60 Hz. Single phase. | 2 |
|  | Vortex Mixer: Variable analog speed control of 100 to 3200 rpm (±25 rpm) with shaker. Rubber feet and a heavy base ensure that the unit won't "walk away" during use. A cup head and 3" dimpled platform. Touch and Auto modes. | 3 |
|  | Ice Flake Maker: Freezing system: Drum rotation freezing jet system, Power: 220V (240V) / 50~60HZ, Production capability: 200- 250Kg, Produce Condition: snow Refrigerant: R-22, or R-404A Machine. | 1 |
|  | DO Meter: DO Range 0.00 to 45.00 ppm (mg/L), 0.0 up to 300.0% saturation DO Resolution 0.01 ppm (mg/L); 0.1% saturation DO Accuracy ±1.5% of reading ±1 digit DO Calibration automatic one or two point at 100% saturation (8.26 mg/L) and 0% saturation (0 mg/L). DO Temperature Compensation ATC: 0 to 50 ºC; 32.0 to 122.0 °F Salinity Compenstaion 0 to 40 g/L (with 1 g/L resolution) DO Altitude Compensation -500 to 4000 m (with 100 m resolution) Temperature Range -20.0 to 120.0 ºC, -4.0 to 248.0 °F Temperature Resolution 0.1 °C, 0.1 °F Logging up to 1000 records organized in: Manual log-on-demand (Max. 200 logs), Manual log-on-stability (Max. 200 logs), Interval logging (Max. 600 samples; 100 lots) Connectivity 1 micro USB port for charging and PC connectivity, 1 USB port for storage Environment 0 to 50°C (32 to 122°F), RH max 95% non-condensing Battery Type/Life Built-in rechargeable battery with up to 8 hours of continuous use Power Supply 5 VDC adapter (included) Dimensions 202 x 140 x 12.7mm (8” x 5.5” x 0.5”). | 2 |
|  | pH Meter (Digital Bench Model): pH range 2 to 19.999, resolution 0.01, accuracy ±0.003, mV range ± 1999mV, manual Temperature, comp 32-212°F, epoxy body combination pH electrode, , 4, 7 and 10 pH standards, swing arm stand, power 9V power supply. | 3 |
|  | EC Meter / Conductivity Meter: Digital Bench Top Microprocessor Controlled with parallel temperature indication, Conductivity Range 0.0 us/cm to 1999 ms/cm (0.001μS to 1mS), Accuracy 0.5%±2 digits, Temperature Range -5.0 to +105.0°C, Salinity 0.0-99.9g/l (0.1), TDS 0-1999 mg/l, Reference Temperature 20 or 25°C Selectable with following accessories. a) Conductivity Electrode with built in Temperature Sensor b) Electrode Stand with Base c) Standard solution 50 ml Complete package, operated on 220 Volts. | 3 |
|  | Vacuum Pump: Motor driven pumps completely mounted on rectangular base plate and furnished with V belt, Single stage. S-25 Displacement 25L/Min, Motor H.P. 0.25. Air Displacement 200 Liters / Minutes. Accessories: - a) Moisture Trap. b) Vacuum Gauge with Regulator. c) Non Return Valve. | 3 |
|  | UV Lamp: UV-light wavelength: 312 nm UV-light intensity variable Viewing area size - 11 x 14 cm UV protection shield for safety UV Mini 11 x 14 cm filter size, 4 x 6 W UV-bulbs (312 nm) with intensity switch; including UV-protecting shield. | 1 |
|  | Drying Cabinet: Timed Dry Cycle: up to 4 hours, Dimensions (External): 1800 x 1168 x 610 mm Three temperature options: Normal, Low and Air Fluff Folding, pull out hangers equivalent of more than 50 feet of clothes line Accessories included: Bar for hangers, Glove hanger, Hanger rod on door, stainless steel shoe rack Very quiet: 60 dB(A). | 1 |
|  | Precision/Micro Pipettes (0.5, 10, 20, 100, 200, 1000 µl): Ajustable volume, Complete Set. | 2 |
|  | Freezer -20C: External dimensions (W x D x H) 770 x 830 x 1955 mm. Internal dimensions (W x D x H) 650 x 700 x 1520 mm. Volume 690 liters Net weight (approx)152 kg Performance Cooling performance-30C - at ambient temperature 35C Temperature setting range -18 ~ -35 C Temperature control range -20 ~ -30C Control Controller Micrprocessor Display LED Temperature sensor Thermistor Refrigeration Cooling method Direct Compressor 400W Refrigerant HFC Insulation material PUF Insulation thickness 60 mm Alarms Power failure V-B-R High temperature V-B-R Low temperature V-B-R Filter - Lid/door open V-B Construction Exterior material Painted steel Interior material Painted steel Outer door/lid Inner door/lid - Outer door/lid lock Y Shelves Drawers - Containers / baskets Optional Max. load -per shelf/basket/drawer 50Max. load -total 200 Access port -position left/top -diameter 30 Interior fluorecent lamp - Casters 4 (2 leveling feet) Electrical and Noise level Power supply 230V 50 Hz Single phase Noise level 42 db. | 1 |
|  | Blenders: AGITATION: Vigorous agitation at a range of frequencies and in multiple directions, for thorough homogenization of even the toughest samples. CAPACITY: Process up to 24 samples in 1.5 mL tubes, 12 samples in 5 mL tubes, or up to 8 samples in 50 mL tubes. COOLING: Dry ice cooling system to keep samples at 4°C, within a few degrees. Sample tube oscillation in open air for good convective cooling. QUIET: Sound insulation to keep homogenization quiet. EASY TO USE: Easy speed and time settings. | 2 |
|  | Liquid Nitrogen Flasks: Max Cont. liter: 1 - 10 Holding time days: 5-16 Evaporation rate l/day: 0.2-0.6. | 1 |
|  | Microwave Oven: Digital with steel body, Steel Panel, Size: 36L and 52 L (One each). | 2 |
|  | Microtome with Accessories: Visual display Range of section thickness:1 - 60µm, Hand wheel: Hand wheel Lock: Positive lock at 12 or 6 o'clock position, Knife Holder Base: Displacement: Lateral Displacement Adjustments: North-south Adjustment Clamping Mechanism : Clockwise Microtome Knives of size 100mm - 300mm, Disposable Blade: Can be used with blade holder, Reverse Wheel: Function: Trimming. | 1 |
|  | Incubator: Temperature control: microprocessor temperature controller Temperature range: 05 C above ambient to 80 C Temperature accuracy: +-0.1 C Temperature uniformity: =- 0.5 at 37 C, Circulation: forced filtered air., Chamber material: stainless steel., Programmable upto 99hours and 59 minutes Power requirements: 220VAC, 50/60Hz single phase. Size 400 x 400 x 500, Volume = 80 liter. | 3 |
|  | 4 Ft Laboratory Fume Hood Bench Top Model: Overall dimension: 52Wx54Hx 24D inchs Appr. Working dimension: 47Wx36Hx 22D inchs Appr External construction: Oven baked epoxy painted MS. Fitted with:· Polypropylene working surface and drainage trough. · Internal walls made of corrosion & acid resistant DEXERON board. · Two electric receptacles to work within the Hood. · Inlets for water, gas & Vacuum. · Vapor proof fluorescent light. · All function control panel for convenient handling. · Heat resistant, acid/corrosion proof safety glass sash counterbalanced with corrosion resistant mechanism · Powerful stainless steel exhaust blower system capable to generate air velocity 100 f/m (avg) at working surface. 220V, 50Hz. Single Phase. | 3 |
|  | Hot Plate / Magnetic Stirrer: Plate Size: 180x180 mm, Stirring capacity (H2O): Max. 20 liters Motor Rating: 9/4 W, Speed control resolution: 5 rpm Speed range: 80 ~ 1,500 rpm, Temp. range: 380 max. Set of magnets (variable size) should be included. | 4 |
|  | Pipette Aid Automatic: Automatic, portable, battery-powered pipet pump equipped with a tissue culture nose piece incorporating a dual hydrophilic, hydrophobic 0.8 μm filter to prevent cross-contamination. Run for two hours (pipetting 2,000 times) without recharging. Should be supplied with 4 replacement filters. Safe, convenient instruments for converting pipettes into hand-operated dispensers. Can be used with glass or plastic pipettes ranging in size from 1 to 100 ml. | 1 |
|  | Protein Electrophoresis Apparatus with All Accessories and Power Supply: Protein Vertical Electrophoresis Apparatus: Four 9 x 10 cm Gels (48 Samples), Gel plates with spacers along gel tank (1000 ml buffer capacity) with lid, combs (12 well), power supply 230V 50/60 Hz, electrode assembly. Seamless Injection-Molded Base, Safety Interlock Cover, Allow Use of Most Prepared Polyacrylamide Gels, Corrosion Resistant, All-Platinum Electrodes, Safety Insulated Electrical Leads, Polyacrylamide Gels (12%)-12 gels, Compatible Power supply to run 4 units at atime should be included. | 2 |
|  | Bench Top Orbital Shaker: The digital orbital shaker suited for the most precise sample incubation and gel de-staining. Digitally controlled RPM from 40-350. Continuous run and pre-set run modes. 12" x 12" platform with skid-free sample pad. | 1 |
|  | Gel Drying System: Gel drying system, 220-240 V, include gel dryer, HydroTech vacuum pump, and instructions. Vacuum-based gel dryer with a drying surface of 35 x 45 cm with variable temperature control and three preprogrammed cycles. | 1 |
|  | Eye Wash Station: Secondary Eye Wash Station Size: 14 x 19 x 5", Capacity: 64 oz. | 2 |
|  | Gel Documentation - Imaging System: UV and LED Easy Electrophoresis gel imaging system (Hood) for DNA, RNA and Protein Gels. 3 Illumination modes (trans-UV, trans white, epi-white), Detector CCD, Compatable with most transilluminator and consumer digital camera. With compatible software, computer, TFT, thermal printer and accessories. | 2 |
|  | Trays for Gels: Consisting of a 2" x 2" pocketless plastic tray stored inside a plastic hinged box. Gel coating applied directly to the surface of the plastic tray. Similar to Gel-Boxes, the Gel-Tray products intended for manual applications in devices loaded using a vacuum pen, tweezers or fingers and unloaded using tweezers or fingers. | 3 |
|  | Micro Kjeldhal Apparatus: for protein estimation Micro Kjeldahl Digestion System Digital Microprocessor Controlled. Digestion Block: Six places, simultaneous digestion with Programming through console, increases of Temperature controlled with microprocessor. Temperature range ambient to 420 - 450 ºC, Setting of heating time 120 minutes, automatic Switch off, possibility to program upto 4 Increases of temperature and twice of permanence each of them. Including digestion tube sets. | 1 |
|  | Electrically Heated Muffle Furnace: Max. Operating Temp. 1200 °C, Volume: 15 litres Weight: 55 kgs, Power rating: 3,6 kW Supply voltage: 1 x 230 V, Inner dimensions: 230x340x170mm (WxDxH) Outer dimensions: 480x650x570mm (WxDxH) 40 segments Programmable start time. | 1 |
|  | Liquid Nitrogen Cylinder: Capacity 40 Liters, 22 psi, Cylinder Size 20 Cu ft | 1 |
|  | Sieve Shaker Mechanical with Sieves Set ASTM: Suitable for large quantities of material to be sieved. Sieve can contain up to 60 to 70 kg of material (approx. 30 liters). Sturdy construction and each sieve blocked individually and easily extractable. The sieve shaker can hold 6 sieves and a receiver. Dustpan door (plexiglass). Supplied with receiver pan and sieves. Power: 220 V, 50 Hz, single phase MHK-200 Air jet Sieve, test sieve shaker for particulate size (5um-5000um). | 1 |
|  | Incubator for BOD: Digital temperature display with high-limit warning light. Capacity: 59 standard, 300 mL BOD bottles (3.1 ft³) or one BOD Trak. Internal dimensions are 22.5 x 16.5x 14.5 inches (H x W x D). External dimensions are 38 x 18.5 x 19 inches (H x W x D). Temperature Range: 5-45°C (Accuracy:±0.5°C @ 20°C). | 1 |
|  | Biofilter for Aquarium: Biological filter media for beneficial bacterial growth * Porous ceramic cylinders, increased surface area for better filtration * Perfect for fresh or saltwater canister filters for cleaner aquarium water, for 10, 000 gallons. | 1 |
|  | Gas Flow Meter: Flow rate (0 - 500) sccm for O2, N2, SiH4. | 3 |
|  | Laser Power Meter: with thermal sensor, photodiode sensor, pyroelectric sensor, beam track sensor preferably, GPIB calbes, USB/RS232, Automation interface, power measurement range 10pW - 100kW, energy measurement range 10 pJ - 40 J. Replaceable battery pack and carrying case along with other accessories. | 2 |
|  | LCR Meter: Fequency range 100 Hz –10 MHz with 4 digit resolution, PC connectivity, LAN, USB, GPIB and cables, test fixtures e.g. Dielectric test fixtures, magnetic test fixtures, clip leads etc. | 2 |
|  | Constant Current and Voltage Programable Power Supply: Current range 0 - 5 A, 20 V, resolution 0.1mA, 1mV. | 2 |
|  | Cryostat for Liquid Nitrogen with Sample Holder: Spectification: Temerature range 77 K to 800 K, can accommodate 6 inch wafers, PID Temperature controller , Micro-Manipulated probe system : Multiple probing of wafers, Cryogenic manipulator probe station , Probe can measure electrical, electro-optical, high Z, DC, RF, and microwave properties of materials and devices. Probe station should provide precise 3-axes control of the probe position to within 1 um resolution. Cooling with liquid nitrogen. | 2 |
|  | Temperature Controller: Specification: Temperature range >1600 C, Controller takes an input from a temperature sensor and has an output that is connected to a control element such as a heater, Temperature sensors. | 2 |
|  | Oscilloscope (dual trace 4 channel) with GPIB Interface and Software: For data acquisition, spectrum analyzer frequency range 9 kHz - 200 MHz, sample rate 2.5 GS/sec. | 2 |
|  | Nd YAG laser (solid state laser with first second and third harmonic generations) YAG Laser with Amplifier: Specifications: Energy: 1000 mJ Wavelength: 1.064um Pulse duration: 10 ns Repetition rate: 100 Hz With seconds harmonics generator (Energy 500 mJ at 532 nm) With third harmonic generator (energy 250 mJ at 355 nm) With fourth harmonic generator (Energy 125 mJ at 266 nm) Temperature stability with integrated wavelength Interfacing and fully computer control and software managements. Water / air cooling system. All harmonic generators should be in compact module including the non linear crystal and removable set of dichroic mirror 0.5 mW He. Ne laser for alignments. Terms and Conditions: Warranty should be of two years. Installation of the module should be done by the company. | 1 |
|  | Spectrometer: Spectrometer range: 200-980 nm with 7 channels Resolution: 0.1 nm (FWHM) Detection:CCDs with a combined 14,336 pixels, Frame rate: 500 Hz capability, computer-controlled Integration time:1 ms; variable in free-run mode, Trigger delay: -121 µs to +135 µs in 500 ns steps, user-configured Trigger jitter: ~20 ns Trigger level: TTL not to exceed 5.5 V Computer connection: USB 1.1 or USB 2.0 Operating systems: Win 7/8/10 on desktop or notebook PCs Software: OOILIBS and OOICOR or equivalent Power requirement: 5 volts @ <1 amp, power supply included Input optical fiber: sampling probe containing 7 fibers, 2 meters long with 74-UV collimating lens and a sampling lens, laser option up to 1000 mJ, continuous spectrum measurements and channels synchornization. | 1 |
|  | Delays: Channels 4 independent pulses controlled in position and width. Range 0 to 2000 seconds Resolution 5 ps Accuracy 1 ns + (timebase error × delay) Jitter (rms) Ext. trig. to any output T0 to any output 25 ps + (timebase jitter × delay) 15 ps + (timebase jitter × delay) Trigger delay 85 ns (ext. trigger to T0 output) External Trigger Rate DC to 1/(100 ns + longest delay). Maximum of 10 MHz Threshold ±3.50 VDC Slope Trigger on rising or falling edge Impedance 1 MΩ + 15 pF Internal Rate Generator Trigger modes Continuous, line or single shot Rate 100 µHz to 10 MHz Resolution 1 µHz Accuracy Same as timebase Jitter (rms) <25 ps (10 MHz/N trigger rate) <100 ps (other trigger rates) Burst Generator Trigger to first T0 Range Resolution 0 to 2000 s 5 ps Period between pulses Range Resolution 100 ns to 42.9 s 10 ns Delay cycles per burst 1 to 232 - 1 Outputs (T0, AB, CD, EF and GH) Source impedance 50 Ω Transition time <2 ns Overshoot <100 mV + 10 % of pulse amplitude Offset ±2 V Amplitude 0.5 to 5.0 V (level + offset <6.0 V) Accuracy 100 mV + 5 % of pulse amplitude General Computer interfaces GPIB (IEEE-488.2), RS-232 and Ethernet. All instrument functions can be controlled through the interfaces. Non-volatile memory Nine sets of instrument configurations can be stored and recalled. Power <100 W, 90 to 264 VAC, 47 Hz to 63 Hz Dimensions 8.5" × 3.5" × 13" (WHL)Weight 9 lbs. Warranty One-Two years parts and labor on defects in materials & workmanship. | 1 |
|  | Optics for Laser: Coated lenses, polarizers, plates for phase shift, Prisms, Optics table- rigid, enhanced stiffness, broadband damping for vibration isolation, precision machined matte finish with +/- 0.1 nm flatness over 1m^2, all steel construction inlcuding top, bottom, core and sides, thick steel 5mm, 430 grade work surface, honey comb, magnetic mounts etc. | 1 |
|  | High Temperature Furnace with Compatible Crucible Box Type (1600 °C): Programable multiple ramp and soak control, 0 - 50 C/ min, 15 Litre or above. | 1 |
|  | Ball Mill, Dry and Wet Grinding: Final fineness < 0.1 micrometer, wheel speed > 500 rev/min. | 1 |
|  | Air Sampler (Basic): Dimensions: 210 x 110 x 220 mm (W x D x H). Case : Powder Coated Aluminium housing. Sampling head : 90 mm head: Autoclavable anodised aluminium head, with 380 holes (Ø 1 mm). Plate support : Autoclavable anodised aluminium plate support for 90 mm ±2 mm plates. Sampling volume (max) 1600 liters. Time to sample 10 min. | 1 |
|  | Hydraulic Press: 20 - 150 ton with heating 500 C - 1000 C, accuracy 0.01+/- Mpa, heating rate 10 C/min. | 1 |
|  | Differential Scanning Calorimeter (DSC): Temperature range: -180°C - 725°C , Heating rates: 0.01 K/min - 100 K/min or more precise, Cooling rates: 0.01 K/min - 100 K/min or more precise, Sensor: Heat flux & power compensation, Vacuum: Yes, PC interface: USB with all accessories including: Gas control boxes, range of crucibles made of gold, silver, aluminum, stainless steel (high pressure etc.), standard sealing press, liquid nitrogen cooling unit (-180°C), Capability for selection of intracoolers as alternative to liquid nitrogen cooling. PC system with 8 GB RAM or more, i7 Processor or higher specs, LED monitor, 1 TB hard disk, DSC software should have following capabilities: Temperature Modulated DSC, Repetition measurements with minimum parameter input, Evaluation of current measurement, Curve comparison up to 16 curves or more, curve subtraction, Zoom function, 1. and 2. Derivative, Multiple smoothing functions, Complex peak evaluation, Multipoint calibration for sample temperature, Multipoint calibration for change of enthalpy, Cp calibration for heat flow, Storage and export of evaluations, Program capable of text editing, Export and import of data ASCII, Data export to MS Excel, Signal-steered measuring procedures, Zoom in function Maintenance kit, Complete Training at least for two persons. 5 year warranty and then after Comprehensive Annual Maintenance Contract for 5 years. | 1 |
|  | GC-MS: GCMS Triple Quadropole System fully PC Controlled The triple quadropole GC-MS should have the following features to deliver high performance operation, maximum sensitivity, maximum uptime, and maximum productivity. GC column oven temperature control from ambient temperature +4 oC to 450oC with atleast 16 heating ramps. Negative ramps should be allowed. Ramp rate = 100 C/min or better, RTL (Retention Time Locking) ready, constant flow or pressure. The system should be provided with inlet for wide range of capillary column. The system should have one FID detector All gases flow should be adjustable/controlled by software. Two 5MS capillary columns (30m x 0.25 mm, film thickness 0.25 µm) and one SP-2330 capillary column or equivalent polarity. Quadrapole should be made up of inert materials to have better mass transfer efficiency and should have heating facility. EI source should be made of inert material and programmable to 300oC or better. Electron Energy 10 to 300 eV. Mass range m/z 10 to 1,050. Scan rate (electronic) of 6000 u/s or better. The system should use low cost Nitrogen as collision gas to reduce cost of operation. Sensitivity of system The sensitivity of system should be as followed and demonstrated at site 100 fg or less octafluoronaphtalene m/z 272-241 transition should give area precision - 4%(n=8) Instrument detection limit - 12 fg or better Basic operation of GC-MS-MS should be controlled from the software having provision for automatic tuning. The MS-MS can be vented and pumped down from the interface for easy maintenance. Qualitative analysis software for GC should be provided. Updated NUST and WILEY libraries should be included. ISI marked hydrogen, nitrogen, zero air, helium cylinder with regulators and gas purification system. A 10 KVA online UPS with 1/2 hours back up. Display screen TFT colour screen size more than 15 cm. PC system with 8 GB RAM, i7 Processor, LED monitor, 1 TB,. ACCESSORIES Autosmapler, Ferrule, gastight and non gastight syringes with Teflon tip plunger, syringes, two MS columns compatible with the instrument. Additional Features Spares to be provided for at least five years. Maintenance kit, Complete Training at least for two persons. 5 year warranty and then after Comprehensive Annual Maintenance Contract for 5 years. System should be sound free. | 1 |
|  | Digital Colony Counter: HD CCD color camera, Digital Zoom 28x, Resolution 1280 x 960 pixels, Minimum size of colonies 0.05 mm, Lighting technology Long-life white LEDs / Dark Field, Automatic counting with manual control, Result export as PDF report, JPEG, PNG, BMP, Excel. | 1 |
|  | Staining Rack: Stainless steel 86×695×21 mm. | 1 |
|  | Phase Contrast Microscope with Accessories: PH condenser 1.25 N.A. (for objectives 10x, 20x, 40x, 100x) with Darkfield (dry) spot, With Camera With all accessories, with cover boxes and cleaning solutions. | 1 |
|  | Refrigerator: 27 Cu Ft, 764 Liters, 2°C to 8°C, 5 adjustable shelves,Temperature control, Digital display Hi/Low alarm, Alarm Relay Dry Contacts, Power failure relay. | 2 |
|  | Autoclave: Gross Capacity: 63L 'Classic' Top Loading SterilizerTemperature Range 100 to 135º C, 316-L Interior Chamber Stainless steel, interior chamber dimension 300x710 D mm, with Built in Timer 1 to 250 Minutes and Pressure Gauge 0 to 0.4 mpa, Selected Modes for Liquid and Solids and Agar Process, Complete with stainless Steel Baskets operated on 220 Volts with following Features:- a. Electromechanical Lock System. b. Dual Sensing Interlock mechanism. c. Pulse Sensing System for Air Exhaust. d. Programmable Auto start e. Process Status Display f. Over Pressure Cutoff. Over Temperature Cutoff. Pressure Safety valve. | 2 |
|  | Centrifuge: With continuously variable speed range 500 to 13000 rpm and time 0 to 99 minutes. Pulse mode setting from 1 to 90 seconds. With rotor or accessories. For 230V 50/60Hz single phase supplies. Centrifuge with capacity of 24 tubes of 2ml. | 1 |
|  | Horizontal Electrophoresis Apparatus: Midi Cell with Gel tray and Power supply, Comb (8, 12 and 16 well) Gel tray size 12x12 and 12 x15 cm, buffer tank 400 ml, 230V 50/60 Hz. | 1 |
|  | Inverted Microscope With Camera: Phase Contrast, Brightfield, Differential Interference Contrast, Relief Contrast, Simple Polarized Light and Fluorescence (Blue/Green Excitations). Mercury, Xenon and Halogen lamps. Tilting Binocular and Trinocular. | 2 |
|  | Shaker Digital: Orbital motion, adjustable timer, digital display of speed, protected against overload and intrinsically safe operation Shaking motion: orbital Shaking movement: 20 mm Speed range infinitely adjustable: 30 – 500 rpm Speed display: Digital Dimensions (WxDxH) 320 x 125 x 375 mm Complete with Platform and fixing Clips 2-Nos each 100, 250, 500ml. | 1 |
|  | Microbiological Safety Cabinet Class II: Overall dimension: WxHxD 1300x1350x650 mm App, Working dimension: WxHxD 1150x675x550 mm App Interior Construction: - 100% 304 stainless steel interior work area, - Built-in water and gas inlets - Electric plugs for instrument, - Regulated air supply, - System with blowers, - Double wall body construction for air circulation, - Grated work surface - Unitized one piece interior side and rear wall construction, - UV & Fluorescent lights, - Main and exhaust Hepa filters Efficiency: 99.9995% @0.12micron Exterior Construction: - 100% 18-gauge cold rolled steel prime and furnished in textured polyurethane - Moveable safety glass window. - All function control, -Pressure gauge. - Time elapsed meter. - Exhaust duct. | 2 |
|  | DIAREADER / Microplate Reader with Computer, Software and Accessories: Detection mode: absorbance > Endpoint, kinetic and well area scanning (via computer control) > 6, 12, 24, 48, 96-well plates 30 seconds for 96 wells single wavelength > Wavelength range: 400 – 750nm (340 nm – 750 nm) > 5 positions filter wheel > Light source: Tungsten Halogen. Computer, software and accessories. | 2 |
|  | Microplate Washer: Microplate types: flat, round or V-shaped wells, 384 or 96 well plates or strips, Wash, prime, rinse and aspirate programs, Up to 75 washer programs stroable. | 1 |
|  | Fluorescence Microscope with Trinocular Head: With camera and accessories, Viewing Head: Trinocular Viewing Head Type: Seidentop Head Rotation: 360° Head Inclination: 30°, Sliding Prism: 70 / 30 Split, Interpupillary Adjustment: 55 – 75mm, Dioptric Adjustment: -5 to +5, Eyepiece Magnification: 10X, Widefield Eyepiece Field Ciameter: 18mm, Power: 220 / 230V, 50Hz. Complete with all standard accessories. | 1 |
|  | Laminar Air Flow: Laminar Flow Benches with Containment: External dimensions 1870x845x1450. 99.999% at 0,3 μm particle size according to DOP 99.995% between 0,12 μm and 0,17 μm particle size according to MPPS Fans: Centrifugal high efficiency with speed control 230 V - 50 Hz single phase. Power consumption (Watts) 525. Fluorescent lights (Watts) 3x40. Ultra violet lights 36, Power socket (Amps) 15. | 2 |
|  | Lyophilizer / Freezer Dryer with all Accessories: With vacuum pumps (Shelf area 3.75 sq ft. Shelf temp -45 C to -70 C, 6Lit in 24 hours, 12 Lit Total Capacity, Floor Model or Compact table top unit for effective lyophilisation of specimens with innovative Wireless Shelf Technology. 85° C condenser temerature for specimens containing solvents or with low eutectic points. Ice condenser capacity: 8kg Ice condenser performance: 6kg/24h Ice condenser temperature -80 C Dimension of Base Model 390 x 415 x 555 mm weight 48 kg. | 2 |
|  | Refrigerated Centrifuge Machine, Temperature Control: Max speed 8000 rpm, vessel capacity, adaptors/buckets for blood collection tube, 15 ml tube, 50 ml tube, 400 ml bottle, refregeration upto 4o C 220-240 V 50/60 Hz, Digital Display with temperature and speed control. | 1 |
|  | Thermal Cycler: Thermal Cycler with 96–Deep Well Reaction Module with gradient (30-100 C), Temperature range 4-100oC, Max ramp rate 4 C per second, Temperature accuracy ± 0.5 oC, VGA colour touch screen display, Memory 500 typical programs, Unlimited with usb drive 220-240 V 50/60 Hz. | 1 |
|  | Sonicator: Sonicator with Temperature Control, Frequency 20-25 kHz, Ultrasonic power (>650W), programable from 1 sec to 10 Hrs, Temperature control range 5-100 C, Digital Temperature monitor with sound enclosure with all accessories, Capacity 10-500 ml 220-240 V 50/60 Hz. | 2 |
|  | Ultra Low Temperature Freezer: Capacity litres 728, Temperature: -50 to -85°C, Upright profile Dual cascade refrigeration system using reliable compressors, Door mounted microprocessor controller with digital set and read temperature display at a comfortable viewing level, Additional internal doors, Polyurethane foamed insulation 130mm thick, Two access ports for temperature probes or gassing. | 1 |
|  | Ordinary Light Microscope with All Acessories: with cover boxes and cleansers, Binocular upright microscope with simple polarized light, objectives (4X, 10X, 40X, 100X) with camera attached with filters. | 4 |
|  | Micro Dispenser: Multi Channel (100-1000 ul). | 1 |
|  | Elisa Plate Reader: Wavelength range 400–750 nm Photometric range 0.0–3.5 OD Linearity ≤1.0% from 0.0–2.0 OD; ≤2.0% from 0.0–3.0 OD Accuracy ≤1.0% or 0.010 from 0.000–3.000 OD at 490 nm Precision 1.0% or 0.005 OD from 0.0–2.0 OD; 1.5% from 2.0–3.0 OD Resolution 0.001 OD Filter wheel capacity 8 Plate shaking (3 speeds) Low, mid, high Duration, sec 0–999 Read time 6 sec at single wavelength, 10 sec at dual wavelengths Data output Onboard graphical thermal printer and USB2 interface with PC or Mac data stations Data storage Calender/clock functions; 64 assay protocols Multilanguage support 4 languages, LCD indication supported; printout report supported Dimensions (W x D x H), cm (in) 34.6 x 37.7 x 16.4 (13.6 x 14.8 x 6.5) Weight, kg (lb) 5.5 (12). Computer, software and accessories. | 1 |
|  | Power Lab with Accessories: Up to 16 input channels, Circadian Analysis, Dose Response, Peak Analysis, Blood Pressure, ECG Analysis, Heart Rate Variability Spike Histogram, Metabolic Cardiac Output. | 1 |
|  | Blood Autoanalyzer with Washing Solutions and Accessories: Whole blood mode: 17 parameters with 3-part Differential, Pre-dilute mode: 8 parameters, which includes WBC, RBC, HGB, HCT, MCV, MCH, MCHC, PLT. Computer, software and accessories. | 1 |
|  | Multimeter: To measure, pH, EC, DO and TDS pH Range 0.0 to 14.0 pH pH Resolution 0.1 pH pH Accuracy ±0.1 pH pH Calibration manual, one point Conductivity Range 0.00 to 4.00 mS/cm Conductivity Resolution 0.01 mS/cm Conductivity Accuracy ±2% F.S. TDS Range 0 to 1999 ppm (mg/L) TDS Resolution 1 ppm (mg/L) TDS Accuracy ±2% F.S. Temperature Range 0.0 to 60.0°C / 32.0 to 140.0°F. | 5 |
|  | Sulphur Dioxide (SO2) Meter: (Ultraviolet Fluorescence method) [Nominal Range 0-20 ppm Maximum 100 ppm Sensor Life > 2 years Resolution 0.1 ppm Min. Detectable 0.1 ppm Response Time < 15 sec. Warm-Up Time 60 sec. Operating Temperature -20 °C to 50 °C Pressure Range Atm ± 10% Relative Humidity 15 - 90 % non-condensing Alarm Audible 80db Dimensions H x D x W: 4.75" x 2.5" x 1.5" Weight 170 grams Battery 9 Volt Alkaline Battery]. | 4 |
|  | Oxides of Nitrogen as NO Meter: (Gas Phase Chemiluminescence) 191 °C Sampling system and 0.4µ inlet filter, Cooled PMT detector, 0-4 ppm to 10,000 ppm without NO2 loss. [Detector Type: Heated Reaction Chamber, Cooled Photomultiplier Tube Sensor Gas Sample System: Stainless Steel, 191 degree C filter and flow control components Reaction Chamber: 65 degree C Stainless Steel reaction chamber Ranges: 0-4, 10, 40, 100, 400, 1000, 4000, 10000 ppm NO / NOx, Auto-ranging standard Response Time: 1.5s (T5 to T95) Gas Flow Rate: 0.5 to 3.0 lpm Inlet Pressure: (-5.0) to 7.0 psig, calibration gas: -5 to 15 psig Accuracy and Repeatability: +/- 1% of range or 0.2ppm Temperature Effects: Zero: <+/- 0.5% from 5 to 35 degrees C, Span: <+/- 0.33% per degree C Detector Noise: <+/- 0.5ppm or 0.1% of range Linearity: < 2.0% of point or 1.0% FS NOX Gas Converter Efficiency: Greater than 95%, test with NOXGEN tester Remote Control: Range, calibration and standby mode via contact closure or RS-232 Outputs: 0-10VDC and isolated 4-20mA concentration and range ID voltage level Logic Outputs: Range, calibration, gas path, mode and readiness Alarms: Two configurable concentration alarms, logic outputs and one fault contact set Power: Dual voltage 115/230 VAC +/- 15%, switch-selectable, 600VA max. Environment: Temperature: 5-40 degrees C, 95%RH, non-condensing]. | 4 |
|  | Oxides of Nitrogen as NO2 Meter: (Gas Phase Chemiluminescence) with Dual Gas Filter Correlation IR and Paramagnetic Oxygen sensor to measure combinations of CO, CO2, NO, O2, SO2 and methane [Response Time: Typically, <15s (T5 to T95) Gas Flow Rate: 1.0 to 5.0 lpm Inlet Pressure: (-5.0) to 15 psig, calibration gas: 5 to 30 psig Accuracy and Repeatability: +/- 1.0% of range or 0.2ppm Temperature Effects: Zero: <+/- 0.2% per degree C, Span: <+/- 0.35% per degree C Detector Noise: Gas and Range dependant Linearity: < 0.5% Remote Control: Range, calibration and standby mode via contact closure or RS-232 Outputs: 0-10VDC and isolated 4-20mA concentration and range identifier voltage Logic Outputs: Range, calibration, gas path, mode and readiness Alarms: Two configurable concentration alarms, logic outputs and one fault contact set Power: Dual voltage 115/230 VAC +/- 15%, switch-selectable, 400VA max]. | 4 |
|  | Tropospheric Ozone (O3) Meter: (Non dispersive UV absorption method) Ambient Ozone Sensor (EZ-1X) Light Source up to 5 tungsten lamps with different narrow band interference filters Light Detector silicon photocell Logging 200 samples Connectivity USB Power Supply external 12 VDC power adapter or built-in rechargeable battery Environment 0 to 50°C (32 to 122°F); RH max 95% non-condensing Dimensions 235 x 200 x 110 mm (9.2 x 7.87 x 4.33”). | 4 |
|  | Suspended Particulate Matter (SPM) Meter: High Volume Sampling, Average not less than 1.1m3/ minute. Suspended Particulate Counter. [Measuring valuesTransmission, opacity, extinction, dust concentration Measuring ranges Transmission: 80 ... 100 % to 0 ... 100 %; adjustable; Opacity: 20 ... 0 % to 100 ... 0 %: adjustable Extiction: 0 ... 0.1 to 0 ... 2.5 dust concentration: 0 ... 800 mg/m3 to 0 ... 12 g/m3, depending on measuring distance Measuring accuracy Transmission: ± 0.4 %, Opacity: ± 0.4 % Extinction: ± 0.002 Measuring distance 0.5 ... 2 m, 1.5 ... 8 m, 6 ...15 m, depending on type of transmitter / receiver unit and reflector Gas temperaturemax. 600 °C Ambient temperature -20 ... +50 °C Protection category IP 65 Measured value output 0/2/4 ... 20 mA analog output; second analog output as an option Interfaces RS 232 for parameterization/service with MEPA; CAN interface as an option Status/control signals 3 relay outp]. | 4 |
|  | Respirable Particulate Matter PM10 Meter: b Ray Absorption method [Measuring range: 0.3 to 1.0, 1.0 ~ 2.5, 2.5 to 10 microns (um) Range: 0 ~ 500 ug / m3 counting accuracy of 50% @ 0.3um, 98% @ ≥0.5 um said the quasi-volume : 0.1 liters (L) Response time: ≤10 seconds (s) DC supply voltage: 5.0 volts (V) Maximum Operating Current: 120 milliamperes (mA) Standby current: ≤200 microamps (uA) Operating temperature range: -20 ~ + 50 degrees Celsius (℃) Operating humidity range: 0 ~ 99% RH MTBF: ≥ 5 years (Y) Maximum size: 65 × 42 × 23 mm]. | 4 |
|  | Digital CO Meter: Carbon monoxide CO (Non Dispersive Infra Red NDIR method) [Response Time: Typically, <15s (T5 to T95) Gas Flow Rate: 1.0 to 5.0 lpm Inlet Pressure: (-5.0) to 15 psig, calibration gas: 5 to 30 psig Accuracy and Repeatability: +/- 1.0% of range or 0.2ppm Temperature Effects: Zero: <+/- 0.2% per degree C, Span: <+/- 0.35% per degree C Detector Noise: Gas and Range dependant Linearity: < 0.5% Remote Control: Range, calibration and standby mode via contact closure or RS-232 Outputs: 0-10VDC and isolated 4-20mA concentration and range identifier voltage Logic Outputs: Range, calibration, gas path, mode and readiness Alarms: Two configurable concentration alarms, logic outputs and one fault contact set Power: Dual voltage 115/230 VAC +/- 15%, switch-selectable, 400VA max. Environment: Temperature: 5-30 degrees C, 95% RH, non-condensing]. | 4 |
|  | Portable Airborne Lead & Lead Dust Monitor: Lead Pb (AAS method after sampling) [Air Sample Time 10-1200 minutes Air Sample Volume – 2-6 Liters/minute sample rate Lead Analysis Time – less than 7 minutes Detection Limit – LoQ Less than 0.05 ug/cubic meter Standard Detection Limit, as Shipped Automatic Operation – Up to 20 continuous samples; Extended Operation Option Available Samples between Battery Charges - 30 (AeroLead(R) 1200) Weight – 12.5 pounds Deminsions – 10.5” h x 7.5” w x 13” d Power Requirements – 110 VAC or internal 12 VDC Battery]. | 4 |
|  | Portable CO2 Meter: Carbon Dioxide CO2 (Digital CO2 Monitor) [CO2 Measurement: non-dispersive infrared (NDIR) Sampling Method: Tube sampling 500ml/min. Response Time: < 20 seconds Range: 1% to 100% CO2 Data Logging: 60,000 data points internal memory Logging interval: 5 seconds to 18 hours Sensor Life Expectancy: > 15 years Calibration certificate tag Maintenance Interval: No maintenance required Tubing size: 1/8" 1,000 hour pump - upgrade available to 10,000 hour pump LED Display CO2 in parts-per-million (ppm) for 1% and 5% models CO2 in percentage format (nn.nn%) for 30%, 60%, 100% Logging On/Off Percentage of memory free]. | 4 |
|  | Digital Sound Level Meter: Noise dB (Digital Noise meter) [±2dB accuracy with 0.1dB resolution A+C weighting AC analog output Record Max/Min values over time Utilizes 0.5"(12.7mm) electret condenser microphone]. | 4 |
|  | Ultra Micro Balance: Particulate matter weight in µg , readibility 0.1µg, capacity 5-21 g, Display 4-6 decimal points. | 4 |
|  | Penetrometer Micrometric Vertical Adjustment: Light-alloy base with leveling screws. Stainless steel calibrated spindle (weight 47.5 g) with friction-free movement. Stop and release button. Dial with 360 divisions corresponding to 0.1 mm penetration steps. Automatic zeroing and relevant micrometric screw. Secured to a cast light-alloy base and slides on the base’s upright. Complete with blocking device and device for the micrometric adjustment of displacements. Suitable for bitumen (needle) and grease (cone). Dimensions: 290 x 270 x 530 (h) mm. Weight: 5.5 kg Accessories supplied as standard: Complete unit with standard penetration needle (for bitumen), 50 and 100 g weights and sample containers: - Diameter 55 x 35 mm (up to 20 mm penetration). | 3 |
|  | Pocket Penetrometer: 0 - 4.5 kgf/sq.cm With a special, calibrated, spring, reading taken directly on the penetration plunger. Measures: 0 - 4.5 kgf/sq.cm Dimensions: 19 mm dia. x 152.4 mm long. Weight: 250g. | 5 |
|  | Stack/Flue Gas Analyzer: O2 Oxygen 0 - 21 % CO2 Carbon dioxide 0 - 25 % (calc.) CO Carbon monoxide 0 - 2,000ppm NOx Nitrogen oxide 0 - 1,200ppm HC Hydro carbon 0 - 100 % LEL SO2 Sulfur dioxide 0 - 4,000ppm H2S Hydrogen sulfide 0 - 500ppm Temp.: - Gas -20 to 1,200°C - Air -20 to 120°C, Draft hPa +/- 40 hPa Lambda / excess air 0 - 999% Losses 0 - 99,9 %, Memory, RS 232 Automatic 0- point calibration Integrated self check program ppm-mg-mg (ref O2) usable at all fuels Printer, Case, Gas Pump, Condense Trap, Filter User manual (engl.) Calibration certificate Battery / charger 115/230 V 50/60 Hz, 12V DC power jack Gas probe Type E 250 mm / 2,5 meter 315x185x380 mm. | 2 |
|  | Total Organic Carbon (TOC) Analyzer: Analysis method: Catalytic high temperature combustion at 850°-950°C (max. 1200°C) Oxidation from C to CO2 which is quantitatively determined with a NDIR. Optional detection of the formed NO with NDIR, CLD or EC. Measuring parameters: TC, TOC, DOC, NPOC, TIC, POC optional TNb, Norms and standards: ISO 8245, EPA 415.1, European standard acc. to EN 1484, ENV 12260, European Pharmacopoeia 6.0, vol.1 (20244) 2008, p. 71, U.S. Pharmacopoeia-NF, USP 30 (643) 2007; DIN 38409, Safety standards: CE, IEC 1010, EMV Basic device - 230 V AC, 50 Hz, 1500 VA - Dimensions (W x H x D)/ Weight: 513 mm x 464 mm x 550 mm, approx. 30 kg - Max. furnace temperature: 950 °C - Gas supply: synthetic air (hydrocarbons and CO2 < 1ppm) or oxygen 99.995 % (4.5) - Measuring range: 0 30.000 mg/l C - Injection volume: 100 µl up to 1000 µl - Analyzing time: approx. 3 - 5 min Required for operation of the analyzer - Desktop PC, tower or laptop - from Intel Pentium 4 - minimum 1 GB RAM, 20 GB HDD - CD ROM drive - interfaces: USB 2.0 - operating system: from MS Windows XP, SP2 or higher - VGA, 16 bit, 1024 x 768 resolution - colour monitor 17´´ (or higher) - Windows compatible, graphic-capable printer - Alternative: stand-alone option incl. stand-alone Software and touch screen Control and data evaluation unit Microtower Intel Core 2 Duo E7600, 1066 MHz, 3 MB 2048 MB DDR3 SDRAM DVD Super Multi Burner HDD > 500 GB Graphic/Audio/LAN on board interface : 1x serial external, 6x USB Windows 7, multilingual Keyboard, Mouse TFT flat screen inkjet printer. | 1 |
|  | Atomic Absorption Spectrometer: High spectral throughput optics with automatic drift correction; 0.3m Czerny-Turner mount monochromator; automated drive grating with 1200 grooves/mm; Wide range high sensitivity PMT 185 to 900 nm; adjustable bandpass settings 0.2, 0.6, 1.2 nm and reduced height 0.6 nm slit for GF analysis; Up to 1000Hz HCL data sampling rate; Self reversal background (Quadline deuterium and AC Zeeman) correction for accurate correction; PC, Windows OS, Microsoft Excel & instrument software -Programmable Gas Control -Automatic 6-Lamp Turret -Switchable True Double Beam Optics -D2 Background Correction -Operation Software -Hi Intensity HCL Power Supply -CB 450 Cooling Bath -Universal Autosampler -Flame AA Starter Kit -Autosampler Kit (F/GF/VG) -High Spectral Purity Hollow Cathode Lamps -1000ppm 500ml Standard solution -Nitrous Oxide Burner Head (Allegheny alloy) -Ventilation Kit -Air Compressor -N2O Gas Regulator with heater -Acetylene (C2H2) Gas Regulator. -T-shaped silica cell for Mercury As standard with Hydride | 1 |
|  | Ion Chromatograph: Single column ion chromatograph for anion (F, Cl, NO3, SO4, PO4) determination in water, wastewater and soil solutions. System: Fully autmated with pump, detector, injection valve, autosampler and peak detection software with full packaged; Operating temp: 10-40 °C; Decible level: 53; Gas pressure: 20-100 psi; Detector: 1000 µS; Auto offset: -999 to 999 µS; Display: Liquid crystal display with 16 front pannel button; Conductivity cell: Chemically inert polymer, 316 stainless steel ectrodes with 1.25 µL active flow and 300 PSI pressure. | 1 |
|  | Portable Total Suspended Solids (TSS) Meter: [It shall have a digital display controlled by microprocessor circuitry. All run, programming, and calibration functions shall be accessible without having to open the enclosure. Unit shall be zero calibrated at the factory. Continuously display suspended solids in mg/l, time, and date. Provide an extended temperature, UV treated LCD digital display. Programming and diagnostics are also provided through this display. Provide self-diagnostics for the sensor and analyzer. Analyzer shall have error messages in the operating mode for higher of lower than normal sensor output voltage, temperature input outside the 0-60 degree C range, unstable instrument circuitry; error messages in the calibrate mode for unstable sensor temperature (after a 5 minute wait), unstable S.S. output of the sensor (after a 5 minute wait), and weak sensor output level. Provide data logging of up to fifty readings. Logged data to include at least SS, location, time, and date. Provide RS-232 serial communication port for suspended solids, time, date, and location stamps. Software for the download of stored data to a computer spreadsheet shall be provided. Provide standard back-lit LCD display. The sensor shall be an near-infrared type sensor operating at 880 nanometers that measures the amount of energy absorbed by the solids in the process. The sensor will automatically compensate for ambient light and shall not be damaged by exposure to direct or indirect sunlight. | 3 |
|  | Chemical Oxygen Demand (COD) Meter: Wavelength range: 320 – 1100 nm Technics: Single Beam with optical cuvette failure elimination. Lamp: Tungsten Bandwidth: 4 nm Measurement modes: Concentration, Absorbance, Transmission, Kinetics and Spectra with Absorbance, %Transmission, Multi-Wavelength, Spectra/Scan: Scans in 1 nm steps for freely defined, wavelength range. Cuvettes: 16 mm round, rectangular 10 mm, 20 mm, 50 mm with automatic recognition, no adapter needed. Display: Backlit graphic display for comprehensive data evaluation. Data storage: 1000 values; for spectra + kinetics approx. 4 MB = 100 spectra (300 – 900 nm) and 400 kinetics with 150 values each. | 1 |
|  | Biological Oxygen Demand (BOD) Meter: BOD Sensor Systems consisting of a 6-position Stirring unit complete with 6 BOD Sensors, 6 Alkali holders for absorbing the carbon dioxide and 6 Stirring bars. Technical Data: Voltage: 230 V / 50 Hz Power: 2W Weight: 3 Kg (6.6 lb) Dimension (WxHxD): 432 x 300 x 165 mm. | 1 |
|  | Fully Automatic Hematology Analyzer: Detailed information of Touch Screen blood Hematology analyzersTest principle: electrical resistance for counting and SFT method for HGBTest flow: fully auto and the blood don't need dilute or other process • Quantity technology: ultra sound quantity method, resistance from disturb • Clot clear: high voltage burn and high pressure flush • Parameter: 3 part differentiation of WBC, 20 parameters + 3 histograms • Sample volume: 9.8μl peripheral blood or venous blood • Test channel: one channel • Throughput: 60 samples/hour • Language: English • QC mode: \ -R\ -B\L-J, auto draw and print QC figure • Calibration mode: auto mode and manual mode are selectable, provide original calibrator and control • Reference range type: newborn\children\male\female and universal • Reagent: provide original reagent • Memory: save more than 30000 sample's detail results including histograms • Display: Large colour LCD, show all parameters and histograms at same screen • Input and output: keyboard\mouse\Parallel Port\RS232 serial port, can be connected to computer • Printer: Inner thermal printer,can be connected to external computer and printer • Working Environment: Temperature:18°~35°; Humidity:40%RH-80%RH • Voltage: AC 175 V~265V, Auto adapt according to the input voltage • Dimension: 300mm*400mm*440mm(L*W*H) • Net weight: 17Kg • Performance • Item Precision, CV • WBC ≤2.0 • RBC ≤1.9 • HGB ≤1.9 • MCV ≤0.4 • PLT ≤4.0 • Carry over WBC ≤2,RBC≤1,HGB≤2,PLT≤2 Technical parameters • Sample volume: normal mode: 9.8μl, dilution mode: 20μlAperture diameter: 70μ m • Quantity technology: ultra sound quantity method, resistance from disturb • Throughput: 60 samples/hour • Clot clear: high voltage burn and high pressure flush • Calibration mode: auto mode and manual mode selectable, provide original calibrator and control • QC mode: \ -R\ -B\L-J, auto draw and print QC figure • Display: Large colour LCD, show all parameters and histograms at same screen • Memory: save more than 30000 sample's detail results including histograms • Input and output: keyboard\mouse\Parallel Port\RS232 serial port, can be connected to computer Printer: Inner thermal printer,can be connected to external computer and printer food packing machine. | 1 |
|  | Calorimeter: Automatic Isoperibol Calorimeter with Oxygen Combustion Bomb, Oxygen Regulator, Nitrogen Regulator, Rinse Tank, 500 Firings Kit, Printer and Ribbon Tests Per Hour: 6 – 8, Operation Time Per Test: 1 Minute Precision Classification: 0.1% Class, Oxygen Fill: Automatic Bucket Fill: Automatic, Balance Communication: RS232C Port Printer Communication: RS232C Port, Power Requirement: 230V, 50 Hz. | 1 |
|  | Viscometer: 5-inch Full Color, Touch Screen Display - User Interface, Enhanced Controls - Real Time Trend Indicator, Supports Multiple Languages Displayed Info: - Viscosity (cP or mPa•s), Temperature (°C or °F) - Shear Rate/Stress, % Torque, Speed/Spindle - Step Program Status Enhanced Security - Customizable User Access, Date and Time Stamp File - Password Access, Portable Log-in Settings Built-in Options - Timed Tests, Data Averaging - Programmable QC Limits/Alarms - Customizable Speed/Spindle Lists - Test Based User Instructions - On Screen Data Comparison Accuracy: ±1.0% of range - Displayed with test data Repeatability: ±0.2% Built-in RTD Temperature Probe Viscosity range 1 cp to 6 million cp, Speed Range (.1 to 200 rpm), Complete with set of 4 spindles, Stand & carrying case, operated on 220 Volts. | 2 |
|  | Amylogram/ Amylograph: Mains connection 1x 230 V; 50/60 Hz + N + PE; 2.8 A 115 V; 50/60 Hz + PE; 5.6 A Heating rate Standard: 1,5 °C/min adjustable 0.1…3.0°C/min Sample volume approx. 550 ml Speed Standard 75 min-1, adjustable 0-300 min-1 Dimensions (W x H x D) 490 x 890 x 400 mm Weight approx. 30 kg net. | 1 |
|  | Mixogram: 35 gram mixograph with case Precision-machined aluminum and steel components mounted on aluminum base plate Aluminum mixing bowls with stainless steel pins. Complete unit mounted on a high quality masonite base and enclosed in an attractive wood cabinet. Electrical Requirements (Specifications with case), 110-120V or 220-440V, 50Hz or 60 Hz, AC, 110-120V, 60Hz, AC, strip chart recorder, AC powered interface, Compatible menu-driven program. | 1 |
|  | Farinogram/ Farinograph: Electric farinograph Parameters Sample Volume cutter 's speed1 cutter's speed2 Test Range Power 300g/50g 63r/min 31.5 r/min 0~10Nm 220V,/250W Features: 1. Advanced mechanical design technology, manufacturing technology and microcomputer control technology, accurate data acquisition 2. High subdivision drive stepper motor rotation smooth and precise. White backlit display, high brightness LCD display 3. Can display viscosity, speed, sample liquid temperature, percent, rotor number, and the selection of the rotor speed, current maximum range value 4. Equipment with min-printer to printer out the test result 5. R232 interface Parameter: Measurement range(mPa.s) 100mPa•s—(2×106)mPa•s (range upto 4×106mPa•s) Rotor speed(r/min) 0.5/1/2/5/10/20/50 Measurement error (F•S) ±1%(F•S) Repeatability ±0.5% printer printout Yes PC communication interface RS232 Power Supply AC(220±10%)V, (50±10% )Hz Temperature control range 45℃~200℃ Temperature accuracy ±0.1℃ Cup volume 20ml Ambient temperature 5C-35C Relative humidity less than 80% Accessories: rotor, 300 ℃ high temperature furnace, Computer system. | 1 |
|  | FPLC (Fast Protein Liquid Chromatography): Variable volume purification columns based on ion-echnage, hydrophobic interaction, size exclusion including sample loading, desalting and guard columns. System pumps Pump type: Piston pump, metering type Flow rate setting: 0.001 to 25 ml/min (normal range) 0.001 to 50 ml/min (column packing flow), Pressure range: 0 to 20 MPa (2900 psi), Viscosity range: 0.35 to 10 cP Sample pump Pump type: Piston pump, metering type Flow rate setting: 0.01 to 25 ml/min, Pressure range: 0 to 10 MPa (1450 psi) Viscosity range: 0.7 to 10 cP Detector 1. UV monitor Wavelength range: 190 to 700 nm in steps of 1 nm, up to 3 wavelengths simultaneously Flow cells: 0.5 mm optical path length, 1μl cell volume 2 mm optical path length, 2 μl cell volume 10 mm optical path length, 8 μl cell volume 2. Conductivity monitor Reading range: 0.01 mS/cm to 999.99 mS/cm 3. Temperature monitor Reading range: 0°C to 99°C 4. pH monitor Reading range: 0 to 14 Integrated fraction collector Number of fractions: up to 576 (6 × 96-well deep-well plates) Vessel types: 3, 8, 15, or 50 ml tubes; 24-, 48-, or 96-well deep-well plates Vessel type selection: Automatic recognition Fraction volumes: 0.1 to 50 ml Cooling specification: 6°C to 20°C, temperature control and all relevant accessories. | 1 |
|  | Texture Analyzer: Product Height: 300 mm Dimensions (HxWxD), Net Weight 64x34x44 cm, 24 kg Load Cells/Force Range: 5, 10, 15, 20, 30, 50, 100 kg. Force Resolution: 1 gf (grams-force) Force Accuracy: <0.025% Cn* Speed Range: 0.1-30 mm/sec Speed Accuracy: 0.03% at 5mm/s Force Accuracy: Better than ±0.017 % Position Accuracy: ± 0.02 mm Precision (RSDr**): ≤ 0.5% Power Requirements: 110-240 V, 50 - 60 Hz * Total combined error (Cn) = Maximum output deviation between the calibration curve and the ideal straight line calibration, % full scale ** Relative Standard Deviation within-instrument repeatability Computer Requirements: PC with Windows XP, Vista or Win7 operating system, DirectX 8 compatible 3D Graphics card, 1 RS 232 com port, 1 USB com port, 1.6 GHz CPU, 512 Mb RAM, hard disk space for data files (>100Mb). | 1 |
|  | Fluorescence Detector: Lamp: Vertically mounted, CW, 150 W Ozone-free xenon arc lamp Gratings: 1200 groove/mm blazed at 330 nm (excitation) and 500 nm (emission), plane ruled Automatic self-calibration of all wavelength drives and slits Monochromators: All reflective optics, Czerny-Turner spectrometers Detectors: Emission: R928P photon counting PMT (185-850 nm) and reference photodiode for monitoring lamp output Water Raman S/N: 6,000:1 (RMS method) Slits: Continuously variable from 0 to 30 nm Accuracy: 0.5 nm Repeatability: 0.1 nm Minimum step: 0.0525 nm Integration time: 0.001 to 160 sec Software: FluorEssence Spectral Correction Factors: Included. | 1 |
|  | Ultrasound Assisted Extraction with Accessories: Capacity 1 liter With double chamber, tub of 0.2 to 1 liter Durable tub, stir rod, all in stainless steel. Double chamber for quicker water cooling along with chiller. Quick connecting attachments for optional cooling/heating. Stainless steel container for module control. Tubular cables with quick connecting attachments. Ultrasound control timer 0-30 min. Variable speed mixer with rods and stainless steel attachments. Operating manual Power 200W 230V 50-60 Hz Dimensions of generator 250x150x360, of stative250x330xh780mm Weight without stirring rod Kg.35. | 1 |
|  | UHT Milk Production Unit: • Capacity of material processing : 20 L per hour. • Heating to 150 ℃ at the flow rates stated. • Cooling to below 5 ℃ at the flow rates stated. • heated hot water recirculators, no steam required • PLC controller, 8 inch LCD touch screen,two modes of run: Automatic-mode, and Eye-monitored mode. • Temperature Precision: ±0.5℃ • The times of sterilizing include: 3S,5S,10S,30 • Product system pressures to 10 bar. • SIP (Sterilise In Place) an option for Aseptic sampling. • Built in true CIP (Clean In Place), flow for CIP more than 400 Ltr per hour • Each heating section is individually controlled to enable a wide setting of product temperatures, heated by pressurised hot water recirculators. • 380 volt 3 phase, 50 or 60 Hz. • Power:13 KW The technological flow Material→ hopper→screw pump→preheating tube→(homo.) →sterilizing tube(137~150℃)→water-cooling→(homo.)→(→ice-cold)→aseptic filling Specifications: Machine Material Stainless Steel Power 12KW/380v Size 1580*1060*1800 Rated capacity 20~25L/h Highest temperature sterilization 150°c Minimum temperature cooling 5℃ precision of temperature control ±0.2℃ sterilizing time 3S 5S 10S 30S 300S maximum pressure 8-10 Bar Minimal sample processing 3~5L. | 1 |
|  | Pilot Plant for Juice Concentrates: Type: Fruit Juice Concentrate Plant Voltage:380V Power: 1.5KW Customized: Customized Color: Silver Material: Stainless Steel Capacity: Ajustable Control: Automatic Function: Orange, Kinds of Fruits Package: Wrapper Film Processing capacity 1 ton-20 tons/h. | 1 |
|  | Blanching and Washing Unit: PPT spray type blanching machine Capacity: 0.5-1.2t/h power: 1.85kw steam consumption: 0.5-1.0t/h water consumption: 1.5-2.5t/h sterilizing time: 10-100min (frequency conversion speed adjustment) dimensions: 12150x1200x1650mm (excluding drying section). | 1 |
|  | Semi-Automatic Food Packaging Machine: Printer: Inner thermal printer, can be connected to external computer and printer food packing machine Specifications: 1. Easy to operate 2. High quality and good price 3. CE certification 4. One year warranty 5. Reliable performance Packing material Heat-sealable laminated film, such as: OPP,PE,PVC,PT/PE,OPP/CPP, other special sealable packing film. Sealing Type: Back-sealed (3 sides sealed) 1.Standard Pillow Style Bag 2.Pillow Style Bag With Hole Punch / Euro Slot 3.Gusseted Pillow Style Bag Accessories: 1. Date coder printer 2. Automatic Punching Device 3. Full Stainless Steel (304/201 ect.) 4.Touch screen/Digital Display (Ultra large touch screen) 5. Release the air 6. No empty bag function 7. Automatic feeder 8. Customized sealing patterns and notch 9. Single/Twin/triple/quad jaws 10. Gas device for modified atmosphere packaging (Nitrogen Charging Device, Air filling device) Features: 1. Double motor, double transducer and PLC controller flexible bag length setting and cutting. 2. Human-machine interface, convenient and quick parameter settings. 3. Self diagnosis failure function. 4. High sensitive optical electric color mark tracking system and digital input cut position. 5. Separated PID control system for temperature. 6. Stopping the machine in selected position, no sticking on the knife and no waste of the film. 7. Simple driving system, reliable working, convenient maintenance. 8. All controlled by program, convenient for function adjusting and upgrading. | 1 |
|  | Digital Flame Photometer: Industrial Model digital single channel with unobstructed samples work area with following features - Safety features including Fail safe flame detection and shut off, with low air pressure indicator. - High sensitivity with fine manual control of air, and stable Flame design. - With ease of use resulting from auto ignition and flame optimization. Peak picker for presenting unambiguous readings with selected filter indicator. With built in Sodium, Potassium, Lithium, Calcium and Barium Filters. Range: ppm: all 1.00 to 9.99, 10.0 to 199.9 mmol/L Minimum Detection Level @ 3 x noise Na 0.1 ppm K 0.1 ppm Li 0.1 ppm Ca & P 2 ppm Ba 5-20 ppm Specificity Na, K and Li interference is <0.5% value of element under analysis, Aspiration rate 3 – 6 mL/min, Speed of Analysis/Time to stability 20 seconds, complete with compressor (but without Gas Cylinder), Operated on 220 Volts. | 2 |

Form- I Technical Requirement

(To be filled by the bidder as part of Technical Compliance) No; Item; Compliance Yes/No; Comment; Part No; Reference Brochure is at page #. A; Equipment. 1; Digital Oven Precise Drying. 2; Rotary Evaporator System. 3; UV/Vis Dual Beam Spectrophotometer. 4; FTIR Spectrometer (Bench-Top). 5; Walk-in Cold Room 10'x10'xS10'. 6; Camera Fitted Light Microscope with Monitors and Accessories. 7; Inverted Microscope with Monitors and Accessories. 8; Shaker Incubator. 9; Glass Distillation Unit, Complete with Mountable Accessories. 10; Electronic/Analytical Balance (Four Digit). 11; Shaking Water Bath (Temperature Controlled). 12; Homogenizer. 13; Vortex Mixer. 14; Ice Flake Maker. 15; DO Meter. 16; pH Meter (Digital Bench Model). 17; EC Meter / Conductivity Meter. 18; Vacuum Pump. 19; UV Lamp. 20; Drying Cabinet. 21; Precision/Micro Pipettes (0.5, 10, 20, 100, 200, 1000µl). 22; Freezer -20C. 23; Blenders. 24; Liquid Nitrogen Flasks. 25; Microwave Oven. 26; Microtome with Accessories. 27; Incubator. 28; 4 Ft Laboratory Fume Hood Bench Top Model. 29; Hot Plate / Magnetic Stirrer. 30; Pipette Aid Automatic. 31; Protein Electrophoresis Apparatus with All Accessories and Power Supply. 32; Bench Top Orbital Shaker. 33; Gel Drying System. 34; Eye Wash Station. 35; Gel Documentation - Imaging System. 36; Trays for Gels. 37; Micro Kjeldhal Apparatus. 38; Electrically Heated Muffle Furnace. 39; Liquid Nitrogen Cylinder. 40; Sieve Shaker Mechanical with Sieves Set ASTM. 41; Incubator for BOD. 42; Biofilter for Aquarium. 43; Gas Flow Meter. 44; Laser Power Meter. 45; LCR Meter. 46; Constant Current and Voltage Programable Power Supply. 47; Cryostat for Liquid Nitrogen with Sample Holder. 48; Temperature Controller. 49; Oscilloscope (dual trace 4 channel) with GPIB Interface and Software. 50; Nd YAG laser (solid state laser with first second and third harmonic generations) YAG Laser with Amplifier. 51; Spectrometer. 52; Delays. 53; Optics for Laser. 54; High Temperature Furnace with Compatible Crucible Box Type (1600 °C). 55; Ball Mill, Dry and Wet Grinding. 56; Air Sampler (Basic). 57; Hydraulic Press. 58; Differential Scanning Calorimeter (DSC). 59; GC-MS. 60; Digital Colony Counter. 61; Staining Rack. 62; Phase Contrast Microscope with Accessories. 63; Refrigerator. 64; Autoclave. 65; Centrifuge. 66; Horizontal Electrophoresis Apparatus. 67; Inverted Microscope With Camera. 68; Shaker Digital. 69; Microbiological Safety Cabinet Class II. 70; DIAREADER / Microplate Reader with Computer, Software and Accessories. 71; Microplate Washer. 72; Fluorescence Microscope with Trinocular Head. 73; Laminar Air Flow. 74; Lyophilizer / Freezer Dryer with all Accessories. 75; Refrigerated Centrifuge Machine, Temperature Control. 76; Thermal Cycler. 77; Sonicator. 78; Ultra Low Temperature Freezer. 79; Ordinary Light Microscope with All Acessories:. 80; Micro Dispenser. 81; Elisa Plate Reader. 82; Power Lab with Accessories. 83; Blood Autoanalyzer with Washing Solutions and Accessories. 84; Multimeter. 85; Sulphur Dioxide (SO2) Meter. 86; Oxides of Nitrogen as NO Meter. 87; Oxides of Nitrogen as NO2 Meter. 88; Tropospheric Ozone (O3) Meter. 89; Suspended Particulate Matter (SPM) Meter. 90; Respirable Particulate Matter PM10 Meter. 91; Digital CO Meter. 92; Portable Airborne Lead & Lead Dust Monitor. 93; Portable CO2 Meter. 94; Digital Sound Level Meter. 95; Ultra Micro Balance. 96; Penetrometer Micrometric Vertical Adjustment. 97; Pocket Penetrometer. 98; Stack/Flue Gas Analyzer. 99; Total Organic Carbon (TOC) Analyzer. 100; Atomic Absorption Spectrometer. 101; Ion Chromatograph. 102; Portable Total Suspended Solids (TSS) Meter. 103; Chemical Oxygen Demand (COD) Meter. 104; Biological Oxygen Demand (BOD) Meter. 105; Fully Automatic Hematology Analyzer. 106; Calorimeter: Automatic Isoperibol Calorimeter with Oxygen Combustion Bomb, Oxygen Regulator, Nitrogen Regulator, Rinse Tank, 500 Firings Kit, Printer and Ribbon. 107; Viscometer. 108; Amylogram/ Amylograph. 109; Mixogram. 110; Farinogram/ Farinograph. 111; FPLC (Fast Protein Liquid Chromatography). 112; Texture Analyzer. 113; Fluorescence Detector. 114; Ultrasound Assisted Extraction with Accessories. 115; UHT Milk Production Unit. 116; Pilot Plant for Juice Concentrates. 117; Blanching and Washing Unit. 118; Semi-Automatic Food Packaging Machine. 119; Digital Flame Photometer. B; Implementation Services. C; Warranty Services.

Note:  Equipment’s Price should also be quoted on Firm’s letter head pad.

|  | DetaiDetails of Call Deposit Receipt (CDR) | Signature______________________________ |  |
| --- | --- | --- | --- |
|  | No__No____________ Amount___________ | Name of Bidder_________________________ |  |
|  | Ban Bank_________ Branch_____ City__________ | CNIC No ______________________________ |  |

Name of the Firm & Address _____________________________Ph No___________________

N.T.No. ____________________________Sales Tax No._______________________________

Officer Incharge
Directorate of Procurement and Inventory Control
Allama Iqbal Road Faisalabad, Phone: 041-9201468 & 9201030

### Annex-II

### Integrity Pack Format

DECLARATION OF FEES, COMMISSION AND BROKERAGE ETC.

PAYABLE BY THE SUPPLIERS OF GOODS, SERVICES & WORKS IN

CONTRACTS WORTH RS.10.00 MILLION OR MORE

Contract Number: ________________________ Dated: ________________________

Contract Value:_____________________________

Contract Title: ________________________________

________________ hereby declares that it has not obtained or induced the procurement of any contract, right, interest, privilege or other obligation or benefit from Government of Pakistan or any administrative subdivision or agency thereof or any other entity owned or controlled by it (GoP) through any corrupt business practice.

Without limiting the generality of the foregoing, ____________ represents and warrants that it has fully declared the brokerage, commission, fees etc. paid or payable to anyone and not given or agreed to give and shall not give or agree to give to anyone within or outside Pakistan either directly or indirectly through any natural or juridical person, including its affiliate, agent, associate, broker, consultant, director, promoter, shareholder, sponsor or subsidiary, any commission, gratification, bribe, finder's fee or kickback, whether described as consultation fee or otherwise, with the object of obtaining or inducing the procurement of a contract, right, interest, privilege or other obligation or benefit in whatsoever form from GoP, except that which has been expressly declared pursuant hereto.

________________ certifies that it has made and will make full disclosure of all agreements and arrangements with all persons in respect of or related to the transaction with GoP and has not taken any action or will not take any action to circumvent the above declaration, representation or warranty.

_____________ accepts full responsibility and strict liability for making any false declaration, not making full disclosure, misrepresenting facts or taking any action likely to defeat the purpose of this declaration, representation and warranty. It agrees that any contract, right, interest, privilege or other obligation or benefit obtained or procured as aforesaid shall, without prejudice to any other right and remedies available to GoPunder any law, contract or other instrument, be voidable at the option of GoP.

Notwithstanding any rights and remedies exercised by GoP in this regard, _____________ agrees to indemnify GoP for any loss or damage incurred by it on account of its corrupt business practices and further pay compensation to GoP in an amount equivalent to ten time the sum of any commission, gratification, bribe, finder's fee or kickback given by _________________ as aforesaid for the purpose of obtaining or inducing the procurement of any contract, right, interest, privilege or other obligation or benefit in whatsoever form from GoP.[Company Name]

Annex-IV

Mandatory Requirement for Participating Bidders

| S. No. | Attributes | Reference Page# |
| --- | --- | --- |
|  | Firms with Income Tax Certificate / GST Certificate. |  |
|  | Affidavit (that the firm has not been blacklisted by private, Govt., Semi Govt. and Autonomous Body) |  |
|  | The bidder should have at least 5 years of relevant experience in supplying, installation, commissioning and technical support of scientific equipements HEIs and R&D organizations. |  |
|  | Quoted product must be European Economic Countries /USA/Canada/Japan origin or the products equivalent in quality with aforementioned countries. |  |